# Gear Measurement Uncertainty and Traceability

# Bruce Cox, Bill Rasnick, and Bill Adkins

Oak Ridge Metrology Center at Lockheed Martin Energy Systems, Inc., Oak Ridge, TN

Measurement uncertainty and traceability of gears and gear artifacts are required by the ISO 9001:19941 and ANSI/NCSL Z540-2-19972 standards.  Although these standards are voluntary in the U. S., failure to meet these requirements could result in the loss of income for gear manufacturers.  The requirement for primary-level calibrations with stated uncertainties and traceability has lead to the development of the National Gear Metrology Center.  This facility was established with funds from the Technology Reinvestment Program as a partnership between the American Society of Mechanical Engineers (ASME), the American Gear Manufacturers Association (AGMA), the National Institute of Standards and Technology (NIST), the Oak Ridge Metrology Center (ORMC), and Pennsylvania State University

# INTRODUCTION

This paper describes the general concept of establishing uncertainty by a method developed by NIST known as measurement decomposition.  This method breaks down the measurement task into its basic components.  The method uses reference artifacts to characterize measurement errors and their effect on the final measurement task.   Specific examples of application are given and guidelines for transferring the uncertainty to shop floor measurements are discussed.

# CONCEPT

Conceptually decomposing the task into its more fundamental subtasks such as alignment, translation, and form measurement can initially seem to be a daunting endeavor.  However, most CMM programmers compose their routines from various simpler tasks as needed, so this type of decomposition can be handled fairly well by one competent in that art.

Decomposition must also be handled considering measuring machine geometry effects.  Following are step-by-step guidelines for the decomposition process.

- Analyze the measuring task: break down total task into elemental tasks.

- Model the effect of error in each elemental task on the final measurement(s).

- Trace the plausible causes of each of these errors back to measuring system error.

- Devise tests to measure each individual machine error or the resultant effect on the individual element task.

- Approximate as closely as possible the actual elemental measurement.

- Use reference artifacts with very low uncertainty, such as gage blocks (length

and angle); spheres; optical flats; and ring gages.

- Calculate the cumulative effect of each of these errors using above model.

- Propagate the uncertainty of each subtask down to the effect on the final measurement.

- Assume a distribution (rectangular, triangular, or Gaussian) for the error effects.

- Calculate the standard deviation of the error effects and combine these effects in quadrature including the uncertainty of the reference artifact(s).  (Refer to NIST Technical Note 12973.)

# EXAMPLES

Three examples of this process are described below.  The first is a gear profile artifact that requires calibration of the portion of the artifact that simulates the gear tooth profile.  The second is a tooth alignment artifact (formerly lead) that requires calibration of the portion of the artifact that simulates a helical surface.  The third is a master gear that requires calibration of the involute profile, the tooth alignment, and the index of the gear teeth.

# Gear Profile Artifact

The measurement task starts with measuring shaft datum features and calculating the centers of these features.  An axis is constructed between these centers.  This axis is the reference axis for the gear profile.  By definition, one point on the gear profile is also a reference.  Once a reference system is established, the profile can be measured (Fig. 1).

The instrument chosen for this task was a Leitz 866 PMM, a high-precision measuring machine with a 3-D analog measuring head  (Fig. 2).

The task decomposition was as follows:

- measuring centers of two circles and calculating line between them,

- translating a certain distance away from this line, and

- measuring a series of points along the involute profile and determining the normal deviation of each point from the theoretical.

The corresponding error effect for each element was modeled and is listed below:

- Any error in determining the centers of the reference or datum surfaces will naturally result in an error in locating the reference axis.  This error can have varying influence upon the profile results depending upon the magnitude and the direction of the misalignment (see Fig. 3).  Software routines that compute the effect of given misalignments were developed and verified by comparison with an independently developed CAD model.

- The profile error resulting from the translation error is variable along the surface of the profile artifact.  Near the start of the artifact the surface normals are nearly perpendicular to the translation direction.  Near the end of the profile the surface normal angles are approximately 60( or less.

- The most influential error is obviously any CMM error in measuring the profile, having a direct 1 to 1 effect.

Likely measuring system errors that would cause the above errors were:

- Probe registration or calibration errors that produce mislocations of the relative positions of each probe with respect to the others could produce an error in the alignment surface location.

- Since the artifact is oriented on the CMM with the reference axis parallel to the CMM y-axis and the initial radial line parallel to the CMM x-axis, any linear error in the x direction at the same approximate z height would produce the translational error.

- Probe tip out of roundness, probe size errors, and probe deflection compensation errors in addition to machine positioning errors will overlay directly on to the profile measurements.

For each of these error models the following tests were devised:

- A circular section of a calibrated sphere was measured twice, once with each probe tip that measures the datum surfaces.

- A gage block of the approximate radius from the center axis to the profile was measured in a line parallel to the radial line from the axis to the initial profile point.

- The profile of a circular section of the calibrated sphere was measured with the section being in the same approximate plane of the profile and the surface normals on the sphere covering the range of surface normals on the profile artifact (see Figs. 4-6).

Not all sources of uncertainty can be determined by tests on reference artifacts.  Although beyond the scope of this paper, these sources and their respective tests are listed for the sake of completeness.

- Any x motion in the y-axis could produce an apparent error in the center axis determination.  This potential error source was tested with a reversal method whereby the involute artifact or a surrogate was measured in 2 opposite orientations (rotated 180( around the z-axis).  The effect of this error was then fed into the appropriate error calculation.

- Repeatability of measurements is always a primary source of uncertainty.  Various causes such as temperature, vibration, and any undetermined factors, while not easy to single out, can be tested for total effect.  By running repeated measurements over an extended period (24-48 h), this variation can be established and statistically characterized.

- The quality of the artifact surfaces themselves can influence the variation between measurements.  Not only the feature of interest (in this case the involute profile) but also the datum surfaces must be of sufficient form and surface quality.  In the case being examined, the roundness of the datum axis features proved to be critical.  Consequently, the roundness had to be characterized before further calibration work could continue.

- Although temperature is controlled closely and monitored, there remains an uncertainty in the temperature measurement and the assumed value for the coefficient of thermal expansion (CTE).  These uncertainties were conservatively assumed to be 0.1( C and 1ppm/( C, respectively.  An estimate of these effects on the profile uncertainty was calculated by expanding and contracting the radial component of the profile coordinates to determine the error limits produced by the temperature and CTE effects.

A rectangular distribution instead of Gaussian was assumed, and the standard deviation calculated from the range of the data.  Two exceptions were the cases for the gage block measurement and the repeatability of measurements, where a Gaussian distribution was appropriate.  In all, there were seven components of uncertainty that were calculated as standard deviations and root-sum-squared to provide the overall standard deviation of uncertainty.  Multiplying by 2 gave a final total 2-sigma (95%) uncertainty of ( 0.9 (m.

In order to verify the uncertainty estimate, three independent measurements were made for inter-comparison.  A surrogate artifact was available from a previous experiment.  This artifact is a spatial arrangement of three precision spheres that closely simulates the basic elements of a gear profile artifact (Fig. 7).  Previous measurements provided an informal 2-sigma estimated uncertainty of ( 0.2 (m on the spherical section that approximates the gear profile.  CMM profile measurements on this artifact and subsequent calculations resulted in a 2-sigma uncertainty of ( 0.8 (m.

Using a third method based on traditional manual metrology methods gave the final affirmation to the uncertainty estimate.  This method essentially duplicated the NIST method developed and reported in the 1980s.  The NIST method uses a precision rotary stage and linear slide to trace the involute profile (unwound string).  Comparison of 0.6 (m between the CMM method and the traditional method is shown in Figs. 8 and 9.

Measurements on two involute artifacts were also made at NIST.  Although NIST equipment is similar, the analysis application software was developed independently, different personnel operate the equipment, and holding fixtures are different.  All NIST and ORMC measurements (on the total mean rise of the profile) were within 0.5 (m see Fig 8.

The consistency of all the inter-comparisons indicated that the uncertainty estimate was valid and perhaps a bit conservative.  After a review of our data and methods, we published a 0.9 (m uncertainty as our calibration service uncertainty.

Gear Tooth Alignment Artifact

A simple description of this artifact would be a helical surface mounted on a shaft with/without mounting centers (Fig. 11).  These artifacts compliment involute profile artifacts and together define the parameters required to inspect a helical gear surface.

The measurement task begins the same way as involute profile artifacts.  The datum features are the centers of the mounting shafts and an axis is constructed between these centers.  By definition, one point on the helical surface is also a reference.

Task decomposition was as follows:

- measuring the centers of two circles and calculating a line between them,

- translating a certain distance away from this line, and

- measuring a series of points along the helical surface and determining the deviation of each point from the theoretical.  This measurement constitutes a translation at an angle to the reference axis.

The corresponding error effect for each element was modeled and is listed below:

- As in the involute profile artifact, any error in determining the centers of the datum surfaces will result in an error in locating the reference axis.  This misalignment of the reference axis influences the error on the helical surface.  The effect depends on the magnitude and direction of the misalignment (see Fig. 12).

- The error resulting from the translation perpendicular to the reference axis must be determined over the range of normal vectors for each lead and along the measured facewidth of the artifact.  This error represents the ability of the CMM to maintain a constant radius from the reference axis along the facewidth of the artifact.

- The error resulting from the translation at an angle to the reference axis must also be determined over the range of normal vectors for each lead and along the measured facewidth of the artifact.  This error represents the ability of the CMM to maintain a constant angle from the reference axis along the facewidth of the artifact.

- The final error comes from the measurement of the profile of the helical curve.  This must also be determined for the ranges of normal vectors of each lead.

Likely measuring system errors that would cause the above errors are as follows:

- Probe registration as in the case of involute profile artifacts.

- Any linear error in the CMM z direction at the same approximate x location would produce the translational error perpendicular to the artifact reference axis.

- Linear errors in both y and x at the same approximate z height would produce the translational error at an angle to the artifact reference axis.

- The same system errors, as in the involute artifact, would overlay directly onto the profile measurements.

For each of these error models the following tests were devised:

- A circular section of a calibrated sphere was measured twice, once with each probe tip that measures the datum surfaces.

- A gage block of the approximate radius from the center axis to the profile was measured in a line perpendicular to each lead angle along the facewidth of the artifact.

- An angle block representing the helix angle of each lead was measured in the same orientation of the corresponding lead on the artifact.

- The profile of a circular section of the calibrated sphere was measured with the section being in the same approximate plane of the helical surface and the surface normal vectors on the sphere covering the range of surface normal vectors on all leads of the artifact (see Figs. 13-17).

As in the case of the involute artifact, other error sources and their respective tests are

listed below:

- Any x motion in the y-axis error was tested by the reversal method whereby the tooth alignment artifact was measured in two opposite orientations (rotated 180( around the z-axis

- Repeatability of measurements was determined by running repeated measurements over an extended period and statistically characterizing this variation.

- The allowable error due to the quality of the artifact was determined to be the same as in the involute artifact.

- The allowable error due to temperature effects was determined to be the same as in the involute artifact.

Measurements on one tooth alignment artifact were also made at NIST.  All NIST

and ORMC measurements (on the mean total rise of the lead variation) were within 0.5 (m (see figure 18).

The traditional manual metrology method, using a precision rotary stage and linear slide to trace the lead surface, was also used.

Comparisons between the traditional and CMM methods indicate agreement within 2.0 (m.  The worst case is shown in Figure 19.

The uncertainty tests and the inter-comparisons were consistent.  We published the

following 95% confidence limit uncertainty values:

Infinite Lead 	(0.8 (m

99 Inch Lead 	(0.9 (m

32 Inch Lead 	(1.1 (m

16 Inch Lead 	(1.2 (m

11 Inch Lead 	(1.3 (m

# Master Gear

In the case of a master gear, the uncertainty determination consists of measuring an involute profile artifact, a tooth alignment artifact, and an index artifact in the same orientation as a gear.  The uncertainty of each artifact has been previously determined measurement decomposition and repeatability measurements are made in the new orientation to determine the uncertainty of each feature ( involute profile, tooth alignment, and index).  For the involute profile artifact and the tooth alignment artifact, the reference surfaces are a plane and one shaft center, when orientated on the rotary table.  The artifacts are shown in Fig. 20-22.

A spur master gear will also be used for internal process control purposes and inter-comparison measurements are being conducted with NIST.

# U95 Uncertainty of the Product Gear

NOTE: It is noted, that for process control, it may be necessary to operate inspection instrument in the same environment as manufacturing equipment.  When the temperature varies significantly from 20( C, the inspection instrument must be calibrated at the operating temperature and the calibrated artifact and product gear should be sufficiently normalized at this temperature.  It may be difficult to predict the effect of temperature on inspection equipment due to geometry and coefficient of thermal expansion (CTE) of different materials, so for certification purposes it is required to correct both the instrument and artifact back to 20( C.

# A. Example - Uncertainty of a gear measuring instrument using the direct comparison method and a calibrated involute master

NOTE: Assume that once the uncertainty of the instrument is determined, only one measurement is taken on the product gear or gear artifact.  If more than one measurement is taken, then the value can be averaged and the U95 uncertainty can be divided by the square root of the number of measurements taken.

- Make at least 10 measurement runs on the gear measuring instrument using the calibrated master as the known artifact.   Ten is not a magic number, more would be better for estimating standard deviations but 10 is a minimum.  These runs should be made very carefully in an effort to obtain the best possible results.

- From the (n) runs in step A1, determine the average ((cm) and standard deviation ((cm).  If the gear measuring instrument is not operated at 20( C, you must either correct the calibrated value (v) to the operating temperature or correct the measured value (x) back to 20( C

- From the values in step A2, determine the lack of agreement (Sd(cm)) between the calibrated value and the measured value.  This lack of agreement is due to the variability in the mean of the measurements and the variability of the calibrated value as shown in Fig. 23 and 24.  The variability in the mean cannot be less than the variability in the calibrated value because each measurement of the artifact is limited by the uncertainty of the calibrated value.   A "student t" test could be used, but it has the complication of which "k" value to use when determining the variability of the artifact.

where

Sv	is the standard deviation of the calibrated value usually expressed as

ncm	is the number of measurements in step 2;

nv	is the number of times the calibrated artifact has been calibrated.

- Over a sufficiently long period of calibrations, the uncertainty of the calibrated artifact could be reduced to the mean of the calibrated artifact plus the non-statistical components of the calibrated artifact uncertainty.

- To determine the bias, subtract the calibrated value from the average from step A2.  If the absolute value of the bias is larger than the absolute value of 2Sd(cm), then a significant bias (more than the 95% confidence level) from the calibrated values exists.  Either the gear measuring instrument should be re-calibrated or the values should be corrected for the bias.

- If the bias is corrected or does not exist, then the U95 uncertainty is two times the root sum squares of the calibrated standard deviation plus the standard deviation of the values in step 2.

- If the bias is not corrected, then the most conservative calculation of the U95 uncertainty is to add the absolute value of the bias to the U95cm (corrected) uncertainty in step 5 including the lack of agreement determined in step 3.

- The uncertainty either from step A5 or A6, depending on the bias, is the uncertainty applied to product gears or to gear artifacts.

# Sample comparison of slope of an involute artifact

# Calibrated value = -0.05 (m

U95(cal) = 0.9 (m

Sv = 0.45 (m

Operating temperature = 22.2222( C

| Run | Slope ((m) | Cor. Slope ((m) |
| --- | --- | --- |
| 1 | 1.3 | 1.5 |
| 2 | 1.2 | 1.4 |
| 3 | 1.1 | 1.3 |
| 4 | 1.0 | 1.2 |
| 5 | 1.5 | 1.7 |
| 6 | 1.0 | 1.2 |
| 7 | 1.6 | 1.8 |
| 8 | 1.2 | 1.4 |
| 9 | 1.3 | 1.5 |
| 10 | 1.5 | 1.7 |

# Avg. ((cm) = -1.474 (m

(cm = 0.211 (m

bias(cm) = -1.474 - (-0.05) = -1.424 (m

2Sd(cm) = 0.910 (m

|-1.424| > |0.910| therefore the bias is significant

# B. Example - Calibrate a working master from a calibrated grand master.

NOTE: Assume that more than one measurement is taken so that the value can be averaged and the U95 uncertainty can be divided by the square root of the number of measurements taken.

# Make at least 10 measurement runs on the gear measuring instrument using the calibrated master as the known artifact and the working master as an unknown. Ten is not a magic number, more would be better for estimating standard deviations but 10 is a minimum.  These runs should be made very carefully in an effort to obtain the best possible results.

- From the (n) runs in step B1, determine the average ((wm) and standard deviation ((wm) of the differences in the two masters.  If the gear measuring instrument is not operated at 20( C, you must either correct the calibrated value (v) to the operating temperature or correct the measured value (x) back to 20( C.

- Since the uncertainty of the working master will always be based on several careful measurements, the standard deviation in step B2 can be divided by the square root of the number of measurements.

# From the values in steps B2 and B3, determine the lack of agreement (Sd(wm)) between the calibrated value and the measured value.

# where

# Sv(cm) is the standard deviation of the calibrated master usually expressed as

# The bias is the average of the differences between the two masters.    If the absolute value of the bias is larger than the absolute value of 2Sd(wm), then a significant bias (more than the 95% confidence level) from the calibrated values exists.   Either the gear measuring instrument should be re-calibrated or the values should be corrected for the bias.

- If the bias is corrected or does not exist, then the U95 uncertainty is two times the root sum squares of the calibrated standard deviation (Sv(cm) from step A.7)  plus the standard deviation of the values in step B3.

# If the bias is not corrected, then the most conservative calculation of the U95 uncertainty is to add the absolute value of the bias to the U95(wm) (corrected) uncertainty in step B6 including the lack of agreement determined in step 4.

- The uncertainty either from step B6 or B7, depending on the bias, is the uncertainty applied to working master.

Sample comparison of working and calibrated involute artifacts (slope)

U95(cal) = 2.77 (m (from A.7)

Sv(cm) = 2.77/2 = 1.39 (m

Operating temperature = 20( C

| Run | Slope Difference ((m) |
| --- | --- |
| 1 | 0.5 |
| 2 | 0.4 |
| 3 | 0.3 |
| 4 | 0.2 |
| 5 | 0.7 |
| 6 | 0.2 |
| 7 | 0.8 |
| 8 | 0.4 |
| 9 | 0.5 |
| 10 | 0.7 |

Avg. ((wm) = -0.47 (m

(wm = 0.211 (m

bias(wm) = -0.47 (m

2Sd(wm) = 2.78 (m

|-0.47| < |2.78| therefore the bias is not significant

# C. Example - Calibrate a product involute master from a working involute master.

NOTE: Assume that the calibration of the product artifacts will be based on the results of a single run rather than the average of more than one run.

# Obtain measurements from at least ten runs of a product involute master against a working master.  These runs can be already available from previous efforts or can be made for this study.  The intent is that they represent routine operating conditions for measurements in an effort to estimate the repeatability of the measurement process when no special care is taken in contrast to the example B above.

- From the (n) runs in step C1, determine the average ((pm) and standard deviation ((pm) of the differences in the two masters.  If the gear measuring instrument is not operated at 20( C, you must either correct the calibrated value (v) to the operating temperature or correct the measured value (x) back to 20( C.

# From the values in step C2, determine the lack of agreement (Sd(pm)) between the working master and the product master.

# where

# Sv(wm) is the standard deviation of the working master usually expressed as

# The bias is the average of the differences between the two masters.    If the absolute value of the bias is larger than the absolute value of 2Sd(pm), then a significant bias (more than the 95% confidence  level) from the calibrated values exists.   Either the gear measuring instrument should be re-calibrated or the values should be corrected for the bias.

# If the bias is corrected or does not exist, then the U95 uncertainty is two times the root sum squares of the calibrated standard deviation (Sv(wm) from step B.8)  plus the standard deviation of the values in step C2.

# If the bias is not corrected, then the most conservative calculation of the U95 uncertainty is to add the absolute value of the bias to the U95(pm) (corrected) uncertainty in step C5 including the lack of agreement determined in step 3.

# The uncertainty either from step C5 or C6, depending on the bias, is the uncertainty applied to product master.

# Sample comparison of product and working involute artifacts (slope)

U95(wm) = 2.78 (m (from B.8)

Operating temperature = 20( C

| Run | Slope Difference ((m) |
| --- | --- |
| 1 | 0.25 |
| 2 | 0.2 |
| 3 | 0.15 |
| 4 | 0.1 |
| 5 | 0.45 |
| 6 | 0.1 |
| 7 | 0.4 |
| 8 | 0.2 |
| 9 | 0.25 |
| 10 | 0.45 |

Avg. ((pm) = -0.255 (m

(pm = 0.134 (m

bias(pm) = -0.255 (m

2Sd(pm) = 2.78 (m

|-0.255| < |2.78| therefore the bias is not significant

# Effect of Temperature on Gear Artifacts

The effect of temperature on profile, tooth alignment, pitch, and runout are described.  For profile, the effect of temperature is essentially due to the change in base circle diameter and the corresponding roll length difference.  To calculate the deviation, determine the roll length of the starting artifact and subtract this value from the roll length of the expanded/contracted base radius due to temperature.

# Base radius = 57.15mm

Roll angle = 20(

Temperature difference = 2.5( C

CTE = 0.0000115 ppm

Roll length1 = Rb1( = (57.15mm)(20()((/180) = 19.949113 mm

Rb2= Rb1+((Rb1)((C)(CTE))= 57.15 + ((57.15)(2.5)(0.0000115)) = 57.151643 mm

# Roll length2 = Rb( = (57.151643 mm)(20()((/180) = 19.949687 mm

Deviation = Roll length2 - Roll length1 = 19.949687 - 19.949113 = 0.000574 mm

For tooth alignment, the effect of temperature is to increase the facewidth of the artifact.  To calculate the effect, determine the helix angle at the radius of measurement, and determine the change in facewidth.  The deviation is the sine of this difference in facewidth.

# Measurement radius = 57.15 mm

Tooth alignment = 279.4 mm

Temperature difference = 2.5( C

CTE = 0.0000115 ppm

Facewidth1 = 45.72 mm

Facewidth2 = Facewidth1 + ((Facewidth1)(2.5)(CTE)) = 45.72 + ((45.72)(2.5)(0.0000115)) = 45.721314 mm

Helix angle = arctan(2(R/L) = arctan((2()(57.15)/279.4)) = 52.113842(

Deviation = (Facewidth2 - Facewidth1)(sin(helix angle)) = (45.721314 - 45.72)(sin(52.113842)) = 0.00104 mm

There is no temperature effect on pitch and runout because the part expands/contracts uniformly and these parameters are based on tooth to tooth differences.

CONCLUSIONS

The main driver for determining measurement uncertainty establishing traceability is to meet the customer demand for certification to international standards.  A key link in this chain is primary level calibration directly traceable to NIST.  In this case, a calibrated artifact is not available.  The measurement decomposition method was developed to determine the uncertainty of primary artifacts.  This method can be employed by anyone in the gear industry or they can transfer the uncertainty and traceability from direct comparison to calibrated artifacts.

# ACKNOWLEDGEMENTS

First and foremost, we must acknowledge Dr. Howard Harary of the NIST for not only originating the decomposition concept, but guiding its implementation.  Mark Cowan of M & M Precision Systems Corporation provided a gear checker machine for use in the ORMC and compiled all of the data used in the inter-comparison between ORMC, NIST, and M & M Precision Systems Corp.

# REFERENCES

- ISO 9001:1994, Quality systems - Model for quality assurance in design, development, production and servicing International Organization for Standardization (Geneva, Switzerland).

- U. S. Guide to the Expression of Uncertainty in Measurement, ANSI/NCSL Z540-2-1997, American National Standards Institute/National Conference of Standards Laboratories, Boulder, CO, 1997.

- Guidelines for Evaluating and Expressing the Uncertainty of NIST Measurement  Results, NIST Technical Note 1297, National Institute of Standards and Technology, Gaithersburg, Md., 1994.

# Text boxes

Figure 2

Figure 3 This is a figure

this is a test 1

Figure 1 Involute profile artifact.

Figure 6 Involute artifact measurement decomposition: profiletest.

Figure 2 Leitz 866 PMM coordinate measuring machine.

Figure 3 Error on involute curve due to misalignment.

Figure 5 Involute artifact measurement decomposition: alignment test B and translation test.

Figure 4 Involute artifact measurement decomposition: alignment test A.

Figure 7 Three-ball surrogate involute artifact.

Figure 11 Tooth alignment artifact.

Figure 8 Traditional manual measurement of involute artifact using a Moore M-18 machine.

Figure 9 Comparison between CMM and manual Moore M-18 for involute artifact.

Figure 10 Comparison between ORMC CMMs, NIST CMM and M&M Gear Checker for involute artifact.

Figure 24 Graphical descripton of U95 uncertainty for uncorrected bias.

Figure 23 Graphical descripton of U95 uncertainty for corrected bias.

Figure 22 Index artifact oriented for master gear calibration.

Figure 21 Tooth alignment artifact oriented for master gear calibration.

Figure 20 Involute profile artifact oriented for master gear calibration.

Figure 19 Comparison of traditional and CMM methods for worst case lead.

Figure 18 Comparison of ORMC and NIST tooth alignment measurements for worst case lead.

Figure 17 Tooth alignment artifact measurement decomposition: probe normal profile test.

Figure 16 Tooth alignment artifact measurement decomposition: angle block profile test.

Figure 15 Tooth alignment artifact measurement decomposition: cylinder radius test.

Figure 14 Tooth alignment artifact measurement decomposition: alignment test B, reversal.

Figure 13 Tooth alignment artifact measurement decomposition: alignment test A, probe registration.

Figure 12 Error on helical surface due to misalignment.